## Slide 1: Plasma-assisted chemical vapor deposition of a-C:H:SiOx coatings. Surface modification of VT6 titanium alloy.
A.S. Grenadyorov1, А.А. Solovyev1, K.V. Oskomov1, V.A. Semenov1, S.V. Rabotkin1, M.O. Zhulkov1,2, A.M. Chernyavskiy2, D.A. Sirota1,2, N.A. Karmadonova3, V.V. Malashchenko1,3, L.S. Litvinova3, O.G. Khaziakhmatova3, N.D. Gazatova3, I.A. Khlusov3,4
1Institute of High Current Electronics SB RAS, Tomsk; 2Meshalkin National Medical Research Center, Novosibirsk; 3Immanuel Kant Baltic Federal University, Kaliningrad; 4Siberian State Medical University, Tomsk. Tomsk – 2021
ISI–2021
The XXV International Conference
on Ion-Surface Interactions

## Slide 2
Titanium materials
(Grade 2, Ti-6Al-4V, Ti-6Al-7Nb et al.)
Problem
Disadvantages:
Low wear resistance,
Insufficient biocompatibility.

## Slide 3
Material volume
Surface layer
Biological environment
Biochemical reactions1,2
Ссылки:
[1] Handbook of Antimicrobial Coatings. Editors: A.Tiwari, A. Chaturvedi. Elsevier. – 2018. – P. 596.
[2] Surgical Tools and Medical Devices. Second Edition. Editors: Waqar Ahmed, Mark J. Jackson. Springer International Publishing Switzerland. – 2016. – P. 691.
[3] Surgical Tools and Medical Devices. Second Edition. Editors: Waqar Ahmed, Mark J. Jackson. Springer International Publishing Switzerland, 2016. – P. 691.
[4] Meskinis S. and Tamuleviciene A. / Structure, Properties and Applications of Diamond Like Nanocomposite (SiOx Containing DLC) Films: A Review // Materials science. – 2011. – V. 17. – №4. – P. 358-370.
[5] Feng Wen, Jiaqi Liu, Jianlu Xue / The Studies of Diamond-Like Carbon Films as Biomaterials: Review // Colloid and Surface Science. – 2017. – V. 2. – P. 81-95.
Solution
a-C:H:SiOx films
High mechanical properties;
Corrosion and wear resistance;
High biocompatibility;
Inactivity to the biological environment of a person3-5.

## Slide 4
Aim of work
The purpose of the work was to modify the surface of the VT6 titanium alloy by a-C:H:SiOx films, characterized by high mechanical, tribological, anticorrosive and biomedical properties.
Tasks:
To obtain a series of samples with a-C:H:SiOx film;
To investigate the tribological and mechanical properties of the obtained samples;
To investigate the corrosion resistance in solutions of 0.5M NaCl and PBS;
To investigate the cytotoxicity of a-C:H:SiOx film in relation to human blood leukocytes;
To investigate the proinflammatory response of a-C:H:SiOx films to cytokines and chemokines.

## Slide 5
Analytical Instruments
Nanoindenter (Nanotest 600, GB);
Tribotester (Pin on the Disc and Oscillating TRIBOtester, France);
Potentiostat-galvanostat P-45X (Electrochemical Instruments, Russia);
Continuous-flow fluorimetry (MACS Quant FL7, Germany);
Flow Analyzer (Bio-Plex Protein Assay System, USA).

## Slide 6
Experimental Vacuum Installation

## Slide 7
Results of Optical Emission Spectrometry of the Discharge
Optical emission spectra of Ar/PPMS mixture plasma in 350–500 nm (a), 500–650 nm (b), 650–800 nm wavelength range at different values of discharge power: (1) – 0 W, (2) – 100 W, (3) –200 W, (4) – 300 W.

## Slide 8
Results of Raman Spectroscopy
Raman spectra of a-C:H:SiOx coatings deposited at different RF discharge power and –300 V substrate bias voltage: a – 0 W, b – 100 W, c – 200 W, d – 300 W.

## Slide 9
Results of EDX Analysis, Mechanical and Tribological Properties
PRF, W | ji, mA/cm2 | C, at.% | Si, at.% | C/Si
- | 0.26±0.02 | 86.0±0.15 | 14.0±0.14 | 6.1
100 | 0.40±0.03 | 87.5±0.10 | 12.5±0.11 | 7.0
200 | 0.65±0.05 | 88.8±0.23 | 11.2±0.28 | 7.9
300 | 0.90±0.06 | 90.1±0.15 | 9.9±0.12 | 9.1
Table 1. Elemental analysis (EDX) of a-C:H:SiOx coatings (Bias -300 V)
PRF, W | H, GPa | H/E | H3/E2, MPa | W, mm3/N·m
0 | 14.3±0.05 | 0.118±0.01 | 199±2 | 5.80·10-7±1.34·10-7
100 | 16.5±0.08 | 0.122±0.01 | 249±2 | 4.66·10-7±1.74·10-7
200 | 16.1±0.26 | 0.121±0.02 | 238±10 | 3.83·10-7±1.43·10-7
300 | 16.4±0.13 | 0.126±0.01 | 260±5 | 2.74·10-7±1.24·10-7
Table 2. Mechanical and tribological properties of a-C:H:SiOx coatings (Bias -300 V)

## Slide 10
Results of Mechanical and Tribological Tests
The deposition of a-C:H:SiOx coating improves the mechanical properties, reduces the coefficient of friction and the wear rate.
Substrate groups | H (GPa) | H/E | H3/E2 (MPa) | µ | W (mm3/N·m)
Ti-6Al-4V | 3±0.1 | 0.03±0.01 | 3±0.8 | 0.63±0.05 | 6.5·10-4±2.5·10-5
Ti-6Al-4V + a-C:H:SiOx | 16±0.2 | 0.13±0.01 | 258±6 | 0.08±0.01 | 4.1×10-7±5.9×10-8
Notation: H – hardness, H/E – plasticity index, H3/E2 – resistance to plastic deformation, µ – coefficient friction, W – wear rate. The films thickness was 1.5 microns.
Mechanical test: Oliver and Pharr method at load 20 mN. Each specimen was indented 20 times, and the obtained results were averaged.
Tribological test (ball-on-disk configuration): 25 mm/s sliding speed, ZrO2 ball of diameter 6 mm used as a counter body, 2.9 mm wear scar, 3 N indentation load resulting in a maximum Hertzian contact pressure of 0.8 GPa, and
500 m sliding distance.
Table 3. Mechanical and tribological properties of a-C:H:SiOx films (PRF=300 W, Bias -130 V)

## Slide 11
Results of Corrosion Tests
The deposition of a-C:H:SiOx coating gives the surface anti-corrosion properties.
Substrate groups | jcorr (mA/cm2) | CR (mm/year)* | Ecorr (mV)
0.5M NaCl solution at 22±2oC |  |  | 
Ti-6Al-4V (T1) | 1·10-8 | 8.7·10-5 | -240
Ti-6Al-4V+a-C:H:SiOx (T2) | 2·10-10 | 15·10-7 | 80
PBS solution at 37±3oC |  |  | 
Ti-6Al-4V (T1) | 3·10-8 | 2.6·10-4 | -288
Ti-6Al-4V+a-C:H:SiOx (T2) | 5·10-9 | 4.1·10-5 | 562
Notation: jcorr – corrosion current density, CR – corrosion rate, Ecorr – corrosion potential.
Corrosion tests: three-electrode measuring setup with P-45X potentiostat/galvanostat (Electrochemical Instruments, Russia); a 0.5 M NaCl solution at 22±2°C and a phosphate-buffered saline (8 g NaCl, 0.2 g KCl, 1.44 g Na2HPO4, 0.24 g KH2PO4) at 37±3°C were used as an electrolyte.
*ASTM, G102-89(2004), Standard Practice for Calculation of Corrosion Rates and Related Information From Electrochemical Measurements, ASTM International, West Con-shohocken, PA, 2004.
Table 4. The results of corrosion tests of obtained samples

## Slide 12
Results of cytotoxicity of the obtained samples
Sample | Number of cells, % | 
 | live | dead
Control | 91.3 | 8.7
Ti-6Al-4V | 91.2 | 8.8
Ti-6Al-4V coated by a-C:H:SiOx film | 90.8 | 9.2
Viable cells amount to over 90%. The obtained results satisfy ISO 10993-5-2009 and prove the absence of a-C:H:SiOx film cytotoxicity for human blood mononuclear leukocytes.
The results were obtained at the Immanuel Kant Baltic Federal University (Kaliningrad).
Table 5. Results of cytotoxicity of the obtained samples

## Slide 13
Test groups, n=5 | Surface potential, mV | Contact angle, deg. | Cytokines |  | Chemokines |  | 
 |  |  | IL-17 | TNF-a | IL-8 | RANTES | MCP-1 (MCAF)
C | - | - | 9 (5-17) | 39 (33-69) | 1436 (1055-1725) | 1035 (636-1166) | 164 (162-167)
T1 | 18±10 | 97±5.1 | 21* (18-22) | 209* (183-311) | 1395* (1242-1570) | 433* (363-659) | 1445* (1324-1570)
T2 | -27±15 | 104±0.5 | 13** (4-18) | 159** (31-230) | 1457** (1064-2514) | 707** (427-956) | 329** (162-370)
T3 | -65±25 | 101±1.5 | 10** (10-27) | 80** (76-87) | 1043** (737-1549) | 377** (287-691) | 188** (80-196)
T4 | -200±50 | 77±1.9 | 12** (9-18) | 58** (34-81) | 803** (591-1152) | 220** (170-598) | 168** (125-177)
Table 6. Concentration (pg/ml) of pro-inflammatory cytokines in intercellular fluid at in vitro mononuclear leukocytes (MNL) contact with the titanium Ti-6Al-4V and Ti-6Al-4V coated by a-C:H:SiOx film for 24 h.
Notes: C – Cells without test samples (control), T1 –Cells with Ti-6Al-4V samples, T2 – Cells with Ti-6Al-4V samples and double-sided deposition of 0.3 μm thick a-C:H:SiOx film, T3 – Cells with Ti-6Al-4V samples and double-sided deposition of 0.8 μm thick a-C:H:SiOx film, T4 – Cells with Ti-6Al-4V samples and double-sided deposition of 1.4 μm thick a-C:H:SiOx film.
In vitro determination of cytokine and chemokine secretions

## Slide 14
It is shown that the mechanical and tribological properties of a-C:H:SiOx coatings depend on the RF discharge power and the bias voltage amplitude of the substrate. It is demonstrated that this parameter can be reduced during the coating deposition without a significant degradation of its mechanical and tribological properties due to the better dissociation and ionization of PPMS vapor molecules in inductively excited RF discharge.
The results of mechanical and tribological tests demonstrated that the obtained a-C:H:SiOx films have high hardness, plasticity index, plastic deformation resistance, and have low friction coefficient and wear rate.
The corrosion tests conducted in a physiological 0.5M NaCl and PBS solutions showed that the deposition of the a-C:H:SiOx film on Ti-6Al-4V samples reduces the corrosion current density from 1∙10-8 to 2∙10-10 A/cm2 and from 3∙10-8 to 5∙10-9 A/cm2, respectively, whereas corrosion rate decreases from 8.7∙10-5 to 15∙10-7 mm/year and from 2.6∙10-4 to 4.1∙10-5 mm/year, respectively.
The biocompatibility assessment showed no pro-inflammatory activity of the film, which in many respects determines the local inflammation development after the implantation.
Cytokine and chemokine concentration decreases with increasing a-C:H:SiOx film thickness from 0.3 to 1.4 μm due to an increase in the negative electrostatic potential and surface wettability.
Conclusions

## Slide 15
Financial Support and Acknowledgment
The work was supported by the Russian Science Foundation under grant No. 19-19-00186.
Acknowledgements:
Solovyev A.A., Head of the Laboratory, Ph.D, IHCE SB RAS (Tomsk).
Oskomov K.V., Senior Researcher, Ph.D., IHCE SB RAS (Tomsk).
Khlusov I.A., Doctor of Medical Sciences, Professor of Siberian State Medical University (Tomsk).
Chernyavskiy A.M., Professor, Doctor of Medical Sciences, Director of the Meshalkin National Medical Research Center (Novosibirsk).
Zhulkov M.O., postgraduate student of the Meshalkin National Medical Research Center (Novosibirsk).
Malashchenko V.V., Research Engineer of the Immanuel Kant Baltic Federal University (Kaliningrad).
Litvinova L.S., MD, Director of the Center for Immunology and Cell Biotechnologies of the Immanuel Kant Baltic Federal University (Kaliningrad).

## Slide 16
List of Main Publications
A.S. Grenadyorov, M.O. Zhulkov, А.А. Solovyev, K.V. Oskomov, V.A. Semenov, A.M. Chernyavskiy, D.A. Sirota, N.A. Karmadonova, V.V. Malashchenko, L.S. Litvinova, O.G. Khaziakhmatova, N.D. Gazatova, I.A. Khlusov / Surface characterization and biological assessment of corrosion resistant a-C:H:SiOx PACVD coating for Ti-6Al-4 V alloy // Materials Science & Engineering C. - 2021. – V. 123. – Article ID 112002. Q1
A.S. Grenadyorov, А.А. Solovyev, N.M. Ivanova, M.O. Zhulkov, A.M. Chernyavskiy, V.V. Malashchenko, I.A. Khlusov / Enhancement of the adhesive strength of antithrombogenic and hemocompatible a-C:H:SiOx films to polypropylene // Surface & Coatings Technology. – 2020. V. 399. Article ID 126132. Q1
A.S. Grenadyorov, V.O. Oskirko, А.А. Solovyev, K.V. Oskomov, I.A. Khlusov / Wear and Corrosion Resistance of a-C:H:SiOx Coating on Medical 316L Stainless Steel // Journal of Materials Engineering and Performance. – 2021. – V. 30. – P. 1099-1109. Q2
A.S. Grenadyorov, A.A. Solovyev, K.V. Oskomov, S.A. Onischenko, A.M. Chernyavskiy, M.O. Zhulkov, V.V. Kaichev / Modifying the surface of a titanium alloy with an electron beam and a-C:H:SiOx coating deposition to reduce hemolysis in cardiac assist devices // Surface and Coatings Technology. – 2020. – V. 381. – Article ID 125113. Q1
A.S. Grenadyorov, A.A. Solovyev, K.V. Oskomov, V.O. Oskirko / Effect of the plasma confinement on properties of a-C:H:SiOx films grown by plasma enhanced chemical vapor deposition // Journal of Vacuum Science & Technology A. – 2019. – V. 37. – P. 061512. Q2
A.S. Grenadyorov, А.А. Solovyeva, K.V. Oskomov, V.S. Sypchenko / Thermal stability of a-C:H:SiOx thin films in hydrogen atmosphere // Thin Solid Films. – 2019. – V. 690. – Article ID 137531. Q2
A.S. Grenadyorov, А.А. Solovyev, K.V. Oskomov, S.V. Rabotkin, Yu.I. Elgin, V.S. Sypchenko, N.M. Ivanova / Effect of substrate bias and substrate/plasma generator distance on properties of a-C:H:SiOx films synthesized by plasma-assisted chemical vapor deposition // Thin solid films. – 2019. – V. 669. – P. 253-261. Q2
A.S. Grenadyorov, A.A. Solovyev, K.V. Oskomov, V.S. Sypchenko / Influence of deposition conditions on mechanical properties of a-C:H:SiOx films prepared by plasma-assisted chemical vapor deposition method // Surface and Coatings Technology. – 2018. – V. 349. – P. 547-555. Q1
A.S. Grenadyorov, А.А. Solovyev, V.V. Malashchenko, I.A. Khlusov / Stability of a-C:H:SiOx coating on polypropylene to chemical sterilization // Journal of Applied Polymer Science. 2020. Article ID e49570. Q2

## Slide 17
  Thanks for your attention!